## LEMI_Install_Sheet
Install Sheet for Long-Period Magnetotelluric Instrument
NETWORK: •✮ | EM
Start date: •✮ | 2020-11-23 | End date: •✮ | 2021-03-02
Field Team: •
Comments: | Network Test
STATION: •✮ | KELLY | Arrival time (UTC): •✮ | 2021-01-02 12:34
Latitude (°): •✮ | 34.06 | Longitude (°): •✮ | 106.89 | Elevation (m): •✮ | 1403 | Declination value (°): ✮ | 8.24
[6 empty] | Declination model: •✮ | EMAG2
Orientation - Reference frame: •✮ | Geomagnetic | Orientation - Method: • | GPS
Recorded components: •✮ | E1, E2, E3, E4, Hx, Hy, Hz
Comments | Station Test
EQUIPMENT
DAS S/N: ✮ | 110 | Model: ✮ | LEMI-424 | Sampling rate (Hz): ✮ | 1
SD Card S/N: ✮ | 5463H7 | Size (MB): | 124
Electrode pair 1 | Positive electrode 1 | S/N: ✮ | 1223 | Direction: •✮ | North | E-field line - Color:
| Negative electrode 1 | S/N: ✮ | 2345 | Direction: •✮ | South | E-field line - Color:
Electrode pair 2 | Positive electrode 2 | S/N: ✮ | 1334 | Direction: •✮ | East | E-field line - Color:
| Negative electrode 2 | S/N: ✮ | 7564 | Direction: •✮ | West | E-field line - Color:
Electrode pair 3 • | Positive electrode 3 | S/N: ✮ | 5647 | Direction: •✮ | North | E-field line - Color:
| Negative electrode 3 | S/N: ✮ | 3425 | Direction: •✮ | South | E-field line - Color:
Electrode pair 4 • | Positive electrode 4 | S/N: ✮ | 1276 | Direction: •✮ | East | E-field line - Color:
| Negative electrode 4 | S/N: ✮ | 8765 | Direction: •✮ | West | E-field line - Color:
Electrode specs: •✮ | Borin STELTH 4 - Silver-Silver Chloride
Fluxgate specs: •✮ | LEMI-039
E-Field Box S/N: ✮ | 546GG5 | Fluxgate S/N: ✮ | 110
SURVEY THE SITE
Comments: | Nice weather
1. Find an area with no close power lines or metallic fences.
2. Ensure that the site is large enough to accommodate the installation of all sensors.
3. Check soil composition, orographic features (hill, valley, sun exposure …) and moisture conditions at places where electrodes will be buried. It is important to select similar
places for the two electrodes composing one measuring line. Clayish grounds are recommended for long-term measurements as they keep the moisture longer.
4. Find north and make sure to write down your geographic coordinate system. It is highly encouraged to use the same coordinate system for both electric and magnetic fields.
INSTALL ELECTRODES
Electrode pair 1 - DC Voltage (V): | 0.012 | Contact Resistance (Ω): | 1100 | Azimuth (°): •✮ | 10 | Approximate tilt (°): •✮ | 3
Electrode pair 2 - DC Voltage (V): | 0.011 | Contact Resistance (Ω): | 1200 | Azimuth (°): •✮ | 100 | Approximate tilt (°): •✮ | 10
Electrode pair 3 • - DC Voltage (V): | 0.013 | Contact Resistance (Ω): | 1040 | Azimuth (°): •✮ | 5 | Approximate tilt (°): •✮ | 2
Electrode pair 4 • - DC Voltage (V): | 0.009 | Contact Resistance (Ω): | 1325 | Azimuth (°): •✮ | 95 | Approximate tilt (°): •✮ | 5
Comments: | Electrodes Test
1. Dig electrode holes. For the majority of soils, a hole depth of about 70 - 80 cm is enough to avoid daily thermal variations.
2. Prepare clayish suspension (about 3 liters for each electrode pair depending upon soil moisture condition).
3. Remove matched pair of electrodes (they are marked) from protective transportation container (do not pull with force!) and apply clayish mixture to each electrode plant.
4. Place electrodes in protective sack, put electrodes in holes. For long-term recording, it is recommended to install electrodes “bottom-up”; for short-term, a tilted position suffices.
5. Burry electrodes by aternatively adding dirt and water to the holes. Protect electric lines with plastic conduits or by burying the lines under the surface.
If you have extra E-field lines, arrange lines in a S shape. If coiled up in a closed loop, the extra lines will behave like an air-core loop and respond to magnetic fields, essentially
becoming an induction coil.
6. Measure the azimuth and tilt of each electrode pair in the specified reference frame of the survey.
7. Measure voltage and contact resistance between electrode pairs. It is best to measure the potential difference from "pot to pot" to get a proper potential difference.
A voltage > 0.1 V or Resistance > 2000-3000 Ω indicates poor contact. In this case, you can add more water or clay to the electrodes or move electrodes to a better location.
INSTALL FLUXGATE
Azimuth (°): •✮ | 20 | Approximate tilt (°): •✮ | 1
Comments: | Fluxgate test
1. Dig fluxgate hole (depth ≥ 50 cm). Place sensor in trash bag for protection. Put covered sensor in hole.
2. Level sensor using bubble level.
3. Turn sensor's housing in such a way that the X arrow on the top of the cover points North. Orientation is accomplished by sensor rotation to achieve zero value
of Y component, then X component will point exactly to Geomagnetic North. Positive directions of the magnetic components are: X – to the North, Y – to the East, Z – down.
4. Once the rotation of the Y component is completed (see step 5 under SET UP DAS), cover the hole with soil and, if available, with sun-reflecting film. Bury magnetic lines.
SET UP POWER SYSTEM AND GPS
Battery voltage (V):
Comments:
1. Test power supply voltage and polarity before connection. Power supply voltage must be within 5 - 20 V range.
2. Connect POWER connector to the 12V battery.
The RED marked wire of the power cable has to be wired to the POSITIVE terminal of the battery and the BLUE marked wire has to be wired to the NEGATIVE terminal!
3. Install GPS antenna using the holder.
INSTALL DAS
Electrode pair 1 | Channel number:✮ | E1 | Dipole length (m):✮ | 78
Electrode pair 2 | Channel number:✮ | E2 | Dipole length (m):✮ | 93
Electrode pair 3 • | Channel number:✮ | E3 | Dipole length (m):✮ | 78
Electrode pair 4 • | Channel number:✮ | E4 | Dipole length (m):✮ | 93
Comments: | Test DAS
1. Connect fluxgate sensor to connector "FG SENSOR" on left side pannel.
2. Connect GPS to connector "GPS" on right side pannel.
3. Connect electrode lines to E-Field box. Normal directions of electric lines pairs have to be: 1st line – positive (RED terminal) to the North, negative (BLACK terminal) to the South;
2nd line - positive to the East, negative to the West. The earth binding post must be connected to a metallic earthing electrode that has to be inserted into the ground directly
near the E-Field box. If any of the electric channels are not in use, short-circuit corresponding pairs of terminals and then connect them to the earthing connector.
4. Record electric channel numbers and dipole lengths.
5. Connect E-Field box to connector "EL.LINES".
6. Insert SD card if not already inserted.
7. Connect power cable to the connector "12V DC". With this the DAS is switched on.
SET UP DAS
SD card - Free space (MB): | 123
GPS location on DAS | Latitude (°): ✮ | 34.05 | Longitude (°): ✮ | 106.9 | Elevation (m): ✮ | 1405
Power supply voltage (V): | 12.3 | TE (°C): | 23 | TF (°C): | 24
Acquisition start time (UTC): •✮ | 2021-01-02 13:43
Comments: | Test Set Up
1. Insert SD card and check space left on card.
2. If GPS antenna is properly connected to the station, GPS module will synchronize internal clock ("Waiting for GPS" in display) and an antenna button will eventually appear
on the screen. Press antenna button to see GPS data in synchronization mode. Verify Date, Time, Latitude, Longitue and Elevation.
3. Record battery voltage. Press tools button, then PA button to access that field. Record electronic unit (TE) and fluxgate sensor (TF) temperatures.
4. Edit dipole length. The length is set to 1 m by default. Press tools button, then LX button. Pressing L1, L2, L3 and L4 will allow to edit the length of each electric line.
Press tools button, then DL button to display length of all electric lines.
5. Check magnetic field value for the y component. Get By as close to 50-100 nT by rotating fluxgate and checking DAS screen. Note: there is a delay of roughly 1s in updating the LCD screen value!
6. Select data acquisition mode (FLash/ PC/ PC + Flash/ LCD) by pressing PC button.
7. Press right arrow to start recording. Report start time of data acquisition.
RUN LEMI-424 PROGRAM | It is not recommended to change values of calibration coefficients specified by the Manufacturer!
Comments: | Nothing to report!
1. Connect field laptop to DAS with USB cable
2. Use the trace display in the main window to check the electric and magnetic field values as well as the fluxgate and logger temperatures.
Channels to display can be selected by using the drop-down lists on the main window toolbar.
3. Look for correlations between different fields (example: when the magnetic fields go up, the electric fields go down) and check amplitude of fields.
4. Generate a copy of the calibration coefficients used by the DAS to internally correct the recorded data.
Select the "tool" button in the main window toolbar. From the newly opened "Station" window, Press the "Coefficients" button. In the "Coefficient" window, select the "save def-file" option.
The coefficients will be saved under the current working directory. Rename the outputted *.def file using the following convention: lemi424N0XXX.def where XXX is the serial number of the DAS.
5. Reset data acquisition mode to FL and disconnect field laptop from DAS.
FINALIZE SITE
Make sure all plugs and connections have been replaced. Cover DAS box with tarp and make sure sensors and field lines are covered with dirt. Pick up any trash.
SITE NOTES
Departure Time (UTC): • | 2021-01-02 14:23
Contact/Landowner:
Direction to site:
Comments/Notes:
1. Take pictures of the site.
2. Record directions to site and any other comments and notes.
• Select cell to display help on screen
✮ Required metadata field
Do not change the layout of this spread sheet! Some of the functionalities of the lemi2seed reader depend on the layout of this template.